## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | CDFW
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | Arnold
Water Temp | °C | 9.66 | 9.7 | 10.32 | 10.4 | -0.04 |  | Performed by: |  | Linarez
SpC | µS/cm | 7285 | 7208 | 7434 | 7332 | 0.3432% |  | Verification Sonde ID: |  | CDFW
DO | mg/L | 9.41 | 9.07 | 9.38 | 9.17 | -0.13 |  | Field Date: |  | 2021-12-17
[6 empty] | -1.382% |  | Exchange Time: |  | 10:41
pH | Units |  |  |  |  | 0
Turbidity | NTU | 17.9 | 17.29 | 15.17 | 15.15 | -0.59 |  | Post-Deployment Date:
[6 empty] | -3.296% |  | Post-Deployment Time: |  | 44912
[8 empty] | Performed By: |  | Linarez
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 20.93 | 20.92 | 0.01 | 0.05 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 12890 | 12885 | 0.0388% | 0.38% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.32 | 9.61 | -0.29 | 0.42 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | -3.018% | 4.4% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units |  |  | 0 | 0 |  | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
0
Turbidity | NTU | 0 | -0.12 | 0.12 | 1.15 | Excellent | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 125 | -1 | 4.1% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks:
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today