## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-20
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | TRB
Water Temp | °C | 9.83 | 9.82 | 10.3 | 10.34 | -0.05 |  | Performed by: |  | Poier
SpC | µS/cm | 795 | 798 | 801 | 802 | 0.2516% |  | Verification Sonde ID: |  | THR-10
DO | mg/L | 10.71 | 10.7 | 10.7 | 10.7 | -0.01 |  | Field Date: |  | 2023-01-19
[6 empty] | -0.09337% |  | Exchange Time: |  | 1115
pH | Units | 7.85 | 7.57 | 7.95 | 7.57 | 0.1
Turbidity | NTU | 117.41 | 100.4 | 132.1 | 114.5 | 0.59 |  | Post-Deployment Date: |  | 2023-01-19
[6 empty] | 0.5025% |  | Post-Deployment Time: |  | 15:00
[8 empty] | Performed By: |  | Poier
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 19.67 | 19.67 | 0 | 0.05 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 24820 | 24828 | -0.03222% | 0.28% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.66 | 9.69 | -0.03 | 0.04 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | -0.3096% | 0.4% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 | 7.1 | -0.08 | 0.17 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.04 | 10.09 | -0.05
Turbidity | NTU | 0 | 0.67 | -0.67 | 1.96 | Excellent | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 121.94 | 2.06 | 2.19% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks:
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today