## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-23
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | BRD
Water Temp | °C | 13.01 | 13.05 | 12.86 | 12.91 | -0.01 |  | Performed by: |  | BA
SpC | µS/cm | 6258 | 6294 | 6218 | 6252 | 0.03196% |  | Verification Sonde ID: |  | THR-21
DO | mg/L | 8.35 | 8.29 | 8.26 | 8.21 | -0.01 |  | Field Date: |  | 2023-11-21
[6 empty] | -0.1198% |  | Exchange Time:
pH | Units | 7.54 | 7.47 | 7.6 | 7.46 | 0.07
Turbidity | NTU | 28.93 | 29.35 | 29.45 | 29.72 | 0.15 |  | Post-Deployment Date: |  | 2023-11-22
[6 empty] | 0.5185% |  | Post-Deployment Time: |  | 11:10
[8 empty] | Performed By: |  | EM
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 19.6 | 19.57 | 0.03 | 0.04 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 12890 | 12895 | % | 0% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.58 | 9.65 | -0.07 | 0.08 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | -0.7254% | 0.85% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 | 7.15 | -0.13 | 0.19 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.04 | 10.15 | -0.11
Turbidity | NTU | 0 | 0.3 | -0.3 | 0.8 | Excellent | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 123 | 1 | 1.33% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks:
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today