## Table 1
Discrete: Real Time: |  | Pre-Measurement Calibration Water Quality Instrument Record California Department of Water Resources Instrument ID: THR-10 Performed by (Last Name): Poier Date of Calibration: 04/03/2023
Site/Run to be used on: TRB deploy Time of Calibration (PST): 09:15 Attachment(s)?
| Verifcation Instruments |  |  |  |  | Calibration Standards |  | Lot # |  |  |  | Expiration Date
| NIST Thermometer S/N | 221360763
[6 empty] | Turbidity Std. |  | 21L21430190 |  |  |  | 2022-11-01
| Therm. Cal. Due Date | 2024-03-18
[6 empty] | Conductance Std. |  | 36176 |  |  |  | 2022-10-31
| Turbidimeter S/N | N/A, 6073G beads
[6 empty] | EC Std. |  | 36179 |  |  |  | 2023-10-31
Turbidmeter Cal. Due Date
[6 empty] | 7 pH Std. |  | 2052C |  |  |  | 2024-02-24
| Barometer S/N | CP307754 |  |  |  | 4 pH Std.
| Barometer Cal. Due Date | 2022-09-21 |  |  |  | 10 pH Std. |  | 2056B |  |  |  | 2023-03-07
Calibration (to be completed within 72 hours before measurement) Followed standard operating procedures version:
* Use pH Standard vs. Temperature Table or calculator **DO %Sat = Local barometric pressure / 7.6 ***Use Dissolved Oxygen Lookup Table (DOLT) or calculator
Parameter |  |  | Standard |  | Pre-Cal |  | Post-Cal |  | Additional Info. |  |  |  | Passing Criteria
Specific Conductance (µS/cm) Dry |  |  | 0 |  | 0 | [7 empty] | < 2 µS/cm
Temperature (°C) in Water |  |  | NIST 2Th0e.r7m1ometer |  | 20.62 | [7 empty] | ≤ ± 0.20 °C
Chlorophyll (RFU) in DIW |  |  | 0 |  | -- | [7 empty] | < 0.10 RFU
Chlorophyll (µg/L) in DIW |  |  | 0 |  | -- | [7 empty] | < 0.10 µg/L
Blue Green Algae (RFU) in DIW |  |  | 0 |  | -- | [7 empty] | < 0.10 RFU
Blue Green Algae (µg/L) in DIW |  |  | 0 |  | -- | [7 empty] | < 0.10 µg/L
Turbidity (NTU) in DIW - Low |  |  | 0 |  | 0.08 |  | 0 | [5 empty] | ≤ ± 0.5 NTU
Turbidity (NTU) in Std - High |  |  | 124 |  | 124.27 |  | 124 | [5 empty] | ≤ ± 5%
Specific Conductance (µS/cm) in Std |  |  | 24820 |  | 24808 |  | 24819.8 |  | Cell Constant: 5.19 |  |  |  | ≤ ± 3%
7 pH (units) |  |  | * 7.02 |  | 7 |  | 7.02 |  | -27.2 mV |  |  |  | ≤ ± 0.2 units
4 pH (units) |  |  | * -- |  | -- |  | -- |  | mV |  |  |  | ≤ ± 0.2 units
10 pH (units) |  |  | * 10.05 |  | 10.06 |  | 10.05 |  | -200.1 mV |  |  |  | ≤ ± 0.2 units
Dissolved Oxygen (%sat) |  |  | ** 100.4 |  | 100.1 |  | 100.1 |  | Baro. Pres. (mmHg): 763.2 |  |  |  | ≤ ± 5%
Dissolved Oxygen (mg/L) |  |  | *** 9.46 |  | 9.41 |  | 9.42 |  | Temp. (°C) in bucket: 18.26 |  |  |  | ≤ ± 0.3 mg/L
Dissolved Oxygen (Gain) | [6 empty] | OD1O.0G3ain |  | 0.87-1.25 range ideal |  |  |  | n/a
172.9
Delta slope (pH 7 mV - pH 10 mV): 1. Calibrate depth to 0 feet: |  |  |  | [Ideal range 160-180 mV, replace module if slope ≤ 155 mV] Comments:
Wiper activated during calibration:
2. Verify sufficient battery voltage for use: 3. Verify wiper is parked correctly & working: 4. Verify or set date & time (Time.gov): Day of Deployment: 1. Verify DO %sat ≤ ± 5% (recommended): 2. Verify parameters & reporting: ✔ 3. Sonde started and ready for measurement:
04/04/2023 0815 Updated May 11th, 2019
Date & Time (PST) of Deployment (RTM) / First Measurement (Discrete):
✔
✔
✔
✔
✔
✔
✔