## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-18
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | BRD
Water Temp | °C | 19.58 | 19.54 | 19.76 | 19.68 | 0.04 |  | Performed by: |  | MO
SpC | µS/cm | 2974 | 3023 | 2990 | 3028 | 0.3699% |  | Verification Sonde ID: |  | THR-19
DO | mg/L | 5.4 | 5.56 | 5.36 | 5.48 | 0.04 |  | Field Date: |  | 2023-10-17
[6 empty] | 0.7407% |  | Exchange Time: |  | 11:26
pH | Units | 7.15 | 7.03 | 7.17 | 7.05 | 0
Turbidity | NTU | 21.83 | 23.7 | 19.37 | 21.01 | 0.23 |  | Post-Deployment Date: |  | 2023-10-18
[6 empty] | 1.054% |  | Post-Deployment Time: |  | 17:00
[8 empty] | Performed By: |  | MO
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 20.98 | 20.96 | 0.02 | 0.06 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 12890 | 12827 | 0.4912% | 0.86% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.01 | 9.02 | -0.01 | 0.05 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | -0.1109% | 0.85% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 | 7.1 | -0.08 | 0.07 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.04 | 10.1 | -0.06
Turbidity | NTU | 0 | 0.08 | -0.08 | 0.83 | Excellent | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 125.13 | -1.13 | 1.96% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks:
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today