## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-18
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | WNG
Water Temp | °C | 11 | 11.03 | 11.1 | 11.09 | 0.04 |  | Performed by: |  | KP, MO
SpC | µS/cm | 8431 | 8282 | 8471 | 8314 | 0.09489% |  | Verification Sonde ID: |  | THR-22
DO | mg/L | 7.36 | 7.47 | 7.44 | 7.44 | 0.11 |  | Field Date: |  | 2023-12-12
[6 empty] | 1.495% |  | Exchange Time: |  | 13:15
pH | Units | 7.5 | 7.56 | 7.49 | 7.54 | 0.01
Turbidity | NTU | 29.7 | 29.4 | 28.6 | 27.9 | 0.4 |  | Post-Deployment Date: |  | 2023-12-12
[6 empty] | 1.347% |  | Post-Deployment Time: |  | 15:00
[8 empty] | Performed By: |  | BA
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 18.5 | 18.5 | 0 | 0.04 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 24820 | 25245 | -1.684% | 1.78% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.77 | 9.7 | 0.07 | 0.18 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | 0.7216% | 2.22% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 |  | 0 | 0.01 |  | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.04 |  | 0
Turbidity | NTU | 0 | 0.1 | -0.1 | 1.05 | Excellent | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 122.8 | 1.2 | 2.32% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks: | after changing batteries and restarting KOR and computer- pH sensor continued to read 6.38 with 11.1mv. Will change pH module. Data QAQC in field looked normal and to be expected- suspected broken bulb can visually see crack- either happened on removal from field or during cleaning
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today