## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-12
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | WNG
Water Temp | °C | 10.8 | 10.82 | 10.86 | 10.84 | 0.04 |  | Performed by: |  | MO
SpC | µS/cm | 4309 | 4332 | 4273 | 4277 | 0.4409% |  | Verification Sonde ID: |  | THR-11
DO | mg/L | 7.79 | 7.8 | 7.75 | 7.67 | 0.09 |  | Field Date: |  | 2024-01-18
[6 empty] | 1.155% |  | Exchange Time: |  | 10:25
pH | Units | 7.54 | 7.51 | 7.56 | 7 | 0.53 |  | Performed By:
Turbidity | NTU | 21.95 | 22.03 | 22.61 | 21.73 | 0.96 |  |  |  | 2024-01-18
[6 empty] | 4.374% |  | Post-Deployment Date: |  | 15:45
Chl | RFU | 0.11 | 0.94 |  |  |  |  | Post-Deployment Time: |  | MO
| ug/L | 0.07 | 0.93
PC | RFU | 7.99 | 5.4
| ug/L | 1.97 | 1.35
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating | Excellent | Good | Fair | Poor | Max. Limit
|  | Standard | Sonde |  |  |  | Excellent | Good
Water Temp | °C | 19.59 | 19.54 | 0.05 | 0.09 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 10000 | 10050.1 | -0.4985% | 0.94% | Excellent | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 9.47 | 9.44 | 0.03 | 0.12 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | 0.3178% | 1.47% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 | 7.07 | -0.05 | 0.57 | Fair | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.04 | 10.06 | -0.02
Turbidity | NTU | 0 | 0.63 | -0.63 | 1.84 | Good | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 122.88 | 1.12 | 5.29% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Chl | RFU | 0 | 0 | 0
| ug/L | 0 | -0.04 | 0.04
PC | RFU | 0 | -0.37 | 0.37
| ug/L | 0 | -0.35 | 0.35
Remarks:
Updated 07/01/2024

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2024-07-01
2024-07-01 | 1.05 | Added Chl and PC value lines. No rating calculations are represented.
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today