## Primary
DRIFT ERROR CALCULATION WORKSHEET
CA Department of Water Resources
Division of Intergrated Science & Engineering - Tidal Habitat Restoration (THR)
|  | Dirty Readings |  | Clean Readings |  | Fouling Error |  | Sonde ID: |  | THR-16
|  | Sonde | Verification | Sonde | Verification |  |  | Removed From: |  | LYR
Water Temp | °C | 22.4 | 22.47 | 22.54 | 22.58 | 0.03 |  | Performed by: |  | BA
SpC | µS/cm | 168.5 | 280.5 | 283.4 | 280.5 | 68.19% |  | Verification Sonde ID: |  | THR-24
DO | mg/L | 6.34 | 6.44 | 6.35 | 6.45 | 0 |  | Field Date: |  | 2024-06-11
[6 empty] | 0% |  | Exchange Time: |  | 8:40
pH | Units | 7.67 | 7.69 | 7.73 | 7.68 | 0.07
Turbidity | NTU | 18.99 | 13.72 | 15.87 | 15.69 | -5.09 |  | Post-Deployment Date: |  | 2024-06-11
[6 empty] | -26.8% |  | Post-Deployment Time: |  | 14:18
[8 empty] | Performed By: |  | BA
|  | Post-Deployment |  | Calibration Error | Total Drift Error | Rating
|  | Standard | Sonde |  |  |  | Excellent | Good | Fair | Poor | Max. Limit
Water Temp | °C | 22.74 | 22.78 | -0.04 | 0.07 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | °C
SpC | µS/cm | 1000 | 1002.6 | -0.2593% | 68.45% | Max. Limit | ≤±3% | ±3-10% | ±10-15% | ±15-30% | >±30% | µS/cm
DO | mg/L | 8.63 | 8.59 | 0.04 | 0.04 | Excellent | ≤±0.3 | ±0.3-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | mg/L
|  |  |  | 0.4657% | 0.47% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-20% | or >±20%
pH | Units | 7.02 | 7.02 | 0 | 0.08 | Excellent | ≤±0.2 | ±0.2-0.5 | ±0.5-0.8 | ±0.8-2.0 | >±2.0 | pH units
|  | 10.03 | 10.04 | -0.01
Turbidity | NTU | 0 | 0.05 | -0.05 | 5.57 | Poor | ≤±0.5 | ±0.5-1.0 | ±1.0-1.5 | ±1.5-3.0 | >±3.0 | NTU
|  | 124 | 123.1 | 0.9 | 27.53% |  | or ≤±5% | or ±5-10% | or ±10-15% | or ±15-30% | or >±30%
Remarks:
Updated 11/05/2019

## Information
|  |  | Excellent | Good | Fair | Poor
|  | Water Temp | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | SpC | 0.03 | 0.1 | 0.15 | 0.3 | 0.3+
|  | DO | 0.3 | 0.5 | 0.8 | 2 | 2.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.2 | 0.2+
|  | pH | 0.2 | 0.5 | 0.8 | 2 | 2.0+
|  | Turbidity | 0.5 | 1 | 1.5 | 3 | 3.0+
|  |  | 0.05 | 0.1 | 0.15 | 0.3 | 0.3+
Changelog |  |  | Last Updated: |  | 2019-11-05
2019-11-05 | 1.04 | Added fields for exchange time and post-deployment time
2019-09-27 | 1.03 | Added Maximum Allowable Limit rating
2018-10-11 | 1.02 | Ratings are now only generated when the data are entered completely
2018-07-25 | 1.01 | Added tabs for multiple sondes from one deployment
2018-07-13 | 1 | I made a thing today